SMART PORTABLE PATIENT MONITORING DEVICE

Clinical Knowledge Base for AI Model Training

Version 1.0  |  Sensors: MAX30102 (HR, SpO₂) · LM35 (Temperature)

Parameters: Heart Rate · SpO₂ · Body Temperature · Derived Indices

Clinical Sources Referenced

- AHA/ACC Guidelines 2022–2023 (Cardiovascular)

- WHO Child Growth Standards & Fever Guidelines

- Royal College of Physicians — NEWS2 2017

- GOLD Guidelines 2023 (COPD)

- GINA Guidelines 2023 (Asthma)

- Surviving Sepsis Campaign 2021

- ADA Standards of Medical Care 2023 (Diabetes)

- APLS / PALS Paediatric Life Support 2021

- British Geriatrics Society 2022

- ESC Heart Failure Guidelines 2021

- BTS Pneumonia Guidelines 2023

- ICMR Population Health Standards (India)

# SECTION 1: AGE-STRATIFIED NORMAL VITAL SIGN RANGES

All thresholds below represent clinically validated normal ranges. Values outside these ranges trigger scoring algorithms in Section 3. Age group boundaries follow AAP, AHA, and WHO paediatric and adult classification standards.

Source: AHA 2023 · WHO Child Growth Standards · AAP PALS 2021 · British Geriatrics Society 2022

## 1.1 Heart Rate Normal Ranges (BPM)

| Age Group | Age Range | HR Min (BPM) | HR Max (BPM) | Bradycardia (<) | Tachycardia (>) |
| --- | --- | --- | --- | --- | --- |
| Infant | 0–12 months | 100 | 160 | 100 | 160 |
| Toddler | 1–3 years | 90 | 150 | 90 | 150 |
| Preschool | 3–5 years | 80 | 140 | 80 | 140 |
| School Age | 6–12 years | 70 | 120 | 70 | 120 |
| Adolescent | 13–17 years | 60 | 100 | 60 | 100 |
| Young Adult | 18–35 years | 60 | 100 | 60 | 100 |
| Middle Aged | 36–59 years | 60 | 100 | 60 | 100 |
| Senior | 60–79 years | 55 | 95 | 55 | 95 |
| Elderly | 80+ years | 50 | 90 | 50 | 90 |

## 1.2 SpO₂ Normal Ranges (%)

| Age Group | Age Range | SpO₂ Min (%) | SpO₂ Max (%) | Hypoxia (<) | Critical (<) |
| --- | --- | --- | --- | --- | --- |
| Infant | 0–12 months | 95 | 100 | 92% | 88% |
| Toddler | 1–3 years | 95 | 100 | 92% | 88% |
| School Age | 6–12 years | 96 | 100 | 93% | 90% |
| Adult (all) | 18–59 years | 95 | 100 | 94% | 91% |
| Senior | 60–79 years | 94 | 100 | 92% | 90% |
| Elderly | 80+ years | 93 | 100 | 91% | 88% |

## 1.3 Body Temperature Normal Ranges (°C)

| Age Group | Age Range | Temp Min (°C) | Temp Max (°C) | Fever Threshold | Hypothermia (<) |
| --- | --- | --- | --- | --- | --- |
| Infant | 0–12 months | 36.5 | 37.5 | >38.0°C | <36.0°C |
| Child | 1–12 years | 36.5 | 37.5 | >38.0°C | <36.0°C |
| Adolescent | 13–17 years | 36.1 | 37.2 | >38.0°C | <35.5°C |
| Adult | 18–59 years | 36.1 | 37.2 | >38.0°C | <35.0°C |
| Senior | 60–79 years | 35.8 | 37.0 | >37.8°C | <35.0°C |
| Elderly | 80+ years | 35.5 | 37.0 | >37.5°C | <34.5°C |

Source: WHO Fever Definition · NICE CG160 · AHA 2023 · BGS 2022

# SECTION 2: COMORBIDITY ADJUSTMENTS TO NORMAL RANGES

The following offsets must be applied to Section 1 ranges when a patient has the listed comorbidity. These adjustments prevent false positive alerts in patients with known chronic conditions. IMPORTANT: For COPD patients, use NEWS2 Scale 2 for SpO₂ scoring (see Section 3).

Source: GOLD 2023 · GINA 2023 · ADA Standards of Care 2023 · AHA/ACC 2022 · ESC 2021 · KDIGO CKD Guidelines

| Comorbidity | HR Adjustment | SpO₂ Adjustment | Temp Adjustment | Clinical Reason |
| --- | --- | --- | --- | --- |
| COPD | Max +10 BPM | Baseline 88–92% acceptable | Afebrile common | Chronic hypoxic drive; O2 therapy target 88–92% |
| Heart Disease | Max −10 BPM | Min −1% | No change | Reduced cardiac reserve; lower HR tolerance |
| Diabetes | No change | No change | Fever threshold −0.3°C | Autonomic neuropathy masks fever & tachycardia response |
| Hypertension | Max −5 BPM | No change | No change | Often on beta-blockers suppressing HR response |
| Asthma | Max +10 BPM | Min −2% (92% = moderate; <90% = critical) | Slight ↑ during exacerbation | GINA severity classification based on SpO₂ |
| CKD | No change | Min −2% | Min −0.2°C | Anaemia and fluid shifts alter SpO₂ baseline |
| Obesity | Max +5 BPM | Min −2% | No change | Hypoventilation syndrome reduces baseline SpO₂ |
| Heart Failure | Max −5 BPM | Baseline 90–95% acceptable | Afebrile unless infective | Compensated HF has chronically low SpO₂; avoid aggressive O2 |

## 2.1 Drug Effects on Vital Signs

These drug effects must be considered when interpreting HR and SpO₂ alerts. If medication data is available, apply the following corrections before scoring:

| Drug Class | Effect on HR | Effect on SpO₂ | Clinical Note |
| --- | --- | --- | --- |
| Beta-blockers | ↓ 10–20 BPM | No direct effect | HR alert thresholds must be lowered by 15 BPM in patients on beta-blockers |
| Calcium channel blockers | ↓ 5–15 BPM | No direct effect | Diltiazem/verapamil class reduces HR; amlodipine does not |
| Bronchodilators | ↑ 10–20 BPM | ↑ during treatment | Salbutamol causes tachycardia; do not flag as cardiac event during nebulisation |
| Opioids | ↓ 5–10 BPM | ↓ dose-dependent | Respiratory depression — SpO₂ drop is primary concern, not HR |
| Antipyretics | ↓ 5–10 BPM | ↑ slightly | Paracetamol/ibuprofen reduce fever-driven tachycardia; do not interpret as clinical improvement alone |

Source: BNF 2023 · AHA Drug Interaction Guidelines · Lexicomp Clinical Pharmacology

# SECTION 3: NEWS2 SCORING SYSTEM (3-SENSOR ADAPTED)

The National Early Warning Score 2 (NEWS2) is the NHS-validated gold standard for detecting patient deterioration. The full NEWS2 uses 7 parameters. This implementation uses 3 available parameters (HR, SpO₂, Temperature) and produces a partial NEWS2 score. All outputs must be labelled "Partial NEWS2 — Respiratory Rate absent."

Source: Royal College of Physicians. National Early Warning Score (NEWS) 2. RCP London, 2017. ISBN: 978-1-86016-737-0

## 3.1 Heart Rate Scoring

| Parameter | Score 3 | Score 2 | Score 1 | Score 0 | Score 1 | Score 2 | Score 3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Heart Rate (BPM) | ≤40 | 41–50 | 51–60 | 61–100 | 101–110 | 111–130 | ≥131 |

## 3.2 SpO₂ Scoring (Scale 1 — Non-COPD) and (Scale 2 — COPD)

| Scale | Score 3 | Score 2 | Score 1 | Score 0 | Use When |
| --- | --- | --- | --- | --- | --- |
| Scale 1 (Default) | ≤91% | 92–93% | 94–95% | ≥96% | No COPD diagnosis. Apply to all other patients. |
| Scale 2 (COPD) | ≤83% | 84–85% | 86–87% | 88–92% | COPD diagnosis confirmed. Scores 1–3 apply above 92% (on supplemental O2). |

## 3.3 Temperature Scoring

| Parameter | Score 3 | Score 1 | Score 0 | Score 1 | Score 2 | Note |
| --- | --- | --- | --- | --- | --- | --- |
| Temperature (°C) | ≤35.0 | 35.1–36.0 | 36.1–38.0 | 38.1–39.0 | ≥39.1 | No score 3 on high side for temperature in standard NEWS2 |

## 3.4 NEWS2 Total Score → Clinical Response

| Total Score | Risk Level | Monitoring Frequency | Required Action |
| --- | --- | --- | --- |
| 0 | Low | Minimum 12-hourly | Continue routine monitoring |
| 1–4 | Low–Medium | 4–6 hourly | Inform registered nurse. Increase monitoring frequency. |
| 5–6 | Medium | Hourly minimum | Urgent review by doctor within 30 minutes. Consider HDU. |
| ≥7 | High | Continuous | Emergency response. Immediate physician review. Consider ICU transfer. |
| ≥3 on any single parameter | Urgent — Single Parameter Flag | Immediate | Even if total score is low, a single score of 3 requires urgent clinical review. |

Source: RCP NEWS2 2017 — Table 2, Table 3. Adapted for 3-parameter implementation.

# SECTION 4: DISEASE-SPECIFIC VITAL SIGN SIGNATURES

The following defines the expected vital sign patterns for each clinical scenario in the training dataset. These patterns are derived from published clinical guidelines and represent the typical presentation, not all possible presentations. The AI model must learn that these patterns are probabilistic, not deterministic.

## 4.1 Sepsis / SIRS

Source: Surviving Sepsis Campaign Guidelines 2021 (Singer et al., JAMA 2016, Sepsis-3)

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Tachycardia >90 BPM (SIRS criterion). May be >120 in severe sepsis. Late-stage septic shock may present with bradycardia <60 BPM — this is an ominous sign. |
| SpO₂ | Normal initially. Falls to 94–96% in moderate sepsis. <92% indicates respiratory compromise. <88% = sepsis-induced ARDS. |
| Temperature | >38.3°C (fever) OR <36.0°C (hypothermia). Hypothermia is associated with worse prognosis. Elderly may not mount a fever response. |
| AI Pattern | HR rising + Temp rising simultaneously over 30–60 minutes is the primary AI detection signature. SIRS = HR >90 AND Temp >38.3 or <36.0 BOTH met. |
| Alert Rule | CRITICAL: Both SIRS criteria met (HR >90 + Temp >38.3 or <36) → Sepsis screening positive. Immediate clinical review required. |

## 4.2 Pneumonia

Source: BTS Guidelines for Management of Community Acquired Pneumonia in Adults 2023

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Tachycardia 100–130 BPM. Correlates with fever severity. CURB-65 score uses HR >125 as a severity criterion. |
| SpO₂ | Mild: 94–96%. Moderate: 90–93%. Severe: <90%. Characteristic pattern is persistent low SpO₂ not responding to position change. |
| Temperature | Typically 38.0–40.5°C. High fever (>39.5°C) with tachycardia is classic presentation. Elderly may have lower fever. |
| AI Pattern | Sustained SpO₂ drop (not transient) + Temp >38°C + HR >100 sustained over multiple readings. SpO₂ characteristically does not recover to baseline between readings. |
| Alert Rule | Moderate concern: SpO₂ <94% sustained + Temp >38°C + HR >100 for >15 minutes. |

## 4.3 COPD Exacerbation

Source: GOLD Guidelines 2023 — Global Strategy for Prevention, Diagnosis and Management of COPD

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | HR 95–120 BPM during exacerbation. Baseline HR often already elevated at 80–95 BPM. Tachycardia driven by hypoxia and anxiety. |
| SpO₂ | KEY RULE: Compare against patient personal baseline, NOT population norms. A drop of ≥4% from personal baseline is the GOLD criterion for exacerbation, regardless of absolute value. Target SpO₂ on O2 therapy is 88–92%, NOT the normal 95%+. |
| Temperature | May be afebrile (viral exacerbation) or mildly febrile 37.5–38.5°C (bacterial exacerbation). High fever >39°C suggests concurrent pneumonia. |
| AI Pattern | Sustained SpO₂ drop from personal baseline (≥4%) + HR elevation + increased HR variability. Absence of fever does not exclude COPD exacerbation. |
| Alert Rule | COPD-specific: SpO₂ drop ≥4% from personal baseline → Exacerbation alert regardless of absolute SpO₂ value. |

## 4.4 Asthma Exacerbation

Source: GINA Guidelines 2023 — Global Strategy for Asthma Management and Prevention

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Moderate: 100–120 BPM. Severe: 120–140 BPM. Life-threatening: may show bradycardia as a pre-arrest sign. |
| SpO₂ | Mild: ≥95%. Moderate: 92–95%. Severe: <92%. Life-threatening: <90%. GINA defines severity primarily by SpO₂. |
| Temperature | Typically normal or slightly elevated 37.0–38.0°C. High fever suggests concurrent infection precipitating the attack. |
| AI Pattern | Rapidly deteriorating SpO₂ (faster than COPD exacerbation) + HR spike + normal or low temperature. SpO₂ drop is typically more acute (minutes, not hours) than COPD. |
| Alert Rule | Severe: SpO₂ <92% in asthma patient → Emergency. Life-threatening: SpO₂ <90% → 999/112 immediately. |

## 4.5 Cardiac Event (ACS)

Source: AHA/ACC Guidelines for Management of Patients with Acute Coronary Syndromes 2022

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Highly variable. STEMI: may have bradycardia <60 BPM (vagal response, especially inferior MI) OR tachycardia >100 BPM. Cardiogenic shock: HR >100 with cold peripheries. |
| SpO₂ | Normal in uncomplicated ACS. Drops to 90–95% in left ventricular failure. <90% indicates acute pulmonary oedema or cardiogenic shock. |
| Temperature | Mildly elevated 37.5–38.5°C peri-infarct (inflammatory response). Normalises within 48 hours. |
| AI Pattern | HR irregularity (SDNN collapse) + SpO₂ drop if cardiac output fails. Bradycardia + SpO₂ drop in the context of chest symptoms = highest ACS suspicion. CRITICAL: HR <40 BPM with SpO₂ drop = potential cardiac arrest precursor. |
| Alert Rule | HIGH RISK: HR <40 or >150 with concurrent SpO₂ <94% → Immediate emergency response regardless of temperature. |

## 4.6 Fever (Isolated)

Source: WHO Fever Definition · NICE CG160 Fever in under 5s · Hasday JD, J Appl Physiol 2000

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | HR rises approximately 10 BPM per 1°C rise above normal body temperature. A patient with Temp 39.5°C should have HR approximately 110–120 BPM. Fever with paradoxically low HR (relative bradycardia) is a red flag for typhoid or brucellosis. |
| SpO₂ | Mildly reduced 94–97%. High fever (>40°C) causes peripheral vasodilation and may lower SpO₂ reading due to perfusion changes. |
| Temperature | Low-grade: 37.3–38.0°C. Fever: 38.0–39.0°C. High fever: 39.1–40.0°C. Hyperpyrexia: >40.5°C (medical emergency). |
| AI Pattern | Proportional HR-Temp relationship is the key AI feature. If HR rise is disproportionate to Temp rise, suspect a secondary complication (sepsis, cardiac involvement). |
| Alert Rule | Hyperpyrexia: Temp >40.5°C → Emergency regardless of HR and SpO₂. |

## 4.7 Hypoxia

Source: WHO Pulse Oximetry Training Manual · AARC Clinical Practice Guidelines

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Compensatory tachycardia 100–140 BPM initially. If hypoxia severe and prolonged, bradycardia develops as a terminal sign. |
| SpO₂ | Mild hypoxia: 90–94%. Moderate: 85–90%. Severe: <85%. Critical: <80% (consciousness affected at this level). |
| Temperature | Normal in isolated hypoxia unless infection is the cause. Temperature is not a primary hypoxia indicator. |
| AI Pattern | SpO₂ below age-adjusted threshold + compensatory HR rise. Trend matters: rapidly falling SpO₂ (>2% per minute) is more concerning than stable low SpO₂. |
| Alert Rule | Moderate: SpO₂ <90% → Urgent. Severe: SpO₂ <85% → Emergency. Critical: SpO₂ <80% → Immediate resuscitation protocol. |

## 4.8 Respiratory Distress

Source: PALS AHA 2021 · APLS UK 2021 · Paediatric specific thresholds

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | HR 120–180 BPM as compensatory response to hypoxia. In paediatrics: HR >180 in infant or >160 in child with SpO₂ drop = impending respiratory failure. |
| SpO₂ | Dropping below 92% with high HR = respiratory distress. <85% = respiratory failure. <80% = impending arrest. |
| Temperature | Variable. May be febrile if infection-driven, or normal if mechanical/anatomical cause. |
| AI Pattern | Combination of maximum age-adjusted HR + falling SpO₂. The simultaneous occurrence of tachycardia AND SpO₂ <92% is the core pattern. In paediatrics, this pattern is more acute than in adults. |
| Alert Rule | PAEDIATRIC CRITICAL: SpO₂ <92% + HR at upper age limit → Immediate intervention. Do not wait for further deterioration. |

## 4.9 Hypertension Crisis

Source: JNC-8 Guidelines · AHA/ACC Hypertension Guidelines 2022

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Typically elevated 100–140 BPM. May be lower if patient is on beta-blockers (masked tachycardia). HR alone is not diagnostic. |
| SpO₂ | Normal unless hypertensive emergency causes pulmonary oedema (SpO₂ drops to 88–94%) or stroke (no direct SpO₂ effect). |
| Temperature | Normal or slightly elevated. Temperature is not a primary hypertension indicator. |
| AI Pattern | HR elevation alone without SpO₂ or Temp change. Without BP sensor, hypertension crisis is not directly detectable — flag as "possible" only if patient has known hypertension diagnosis AND shows unexplained tachycardia with no other cause. |
| Alert Rule | LIMITED DETECTION: Cannot confirm without BP. Flag tachycardia of unknown cause in hypertensive patient for clinical review. |

## 4.10 Heart Failure

Source: ESC Guidelines for Diagnosis and Treatment of Acute and Chronic Heart Failure 2021

| Vital Sign | Expected Pattern & Values |
| --- | --- |
| Heart Rate | Resting tachycardia 90–120 BPM is common in decompensated HF. Worsens with minimal exertion. |
| SpO₂ | Chronically reduced 90–95%. Worsens lying flat (orthopnoea) — if device is wearable, positional SpO₂ change is diagnostic. Acute decompensation: SpO₂ <88%. |
| Temperature | Afebrile unless infective exacerbation of HF. Temperature rise in an HF patient suggests concurrent infection. |
| AI Pattern | Persistent low-grade SpO₂ (90–95%) + resting tachycardia + gradual deterioration over hours-days. Acute SpO₂ drop (<88%) + HR >110 = acute decompensation. |
| Alert Rule | SpO₂ <90% in known HF patient → Acute decompensation. Urgent review required. |

# SECTION 5: TEMPORAL DETERIORATION PATTERNS

The AI model must be trained on temporal sequences, not single readings. The following patterns describe how vitals evolve over time before a clinical event. All time windows are approximate and individual variation is significant.

Source: Churpek et al., Chest 2016 · Tarassenko et al., Resuscitation 2006 · McGaughey et al., Cochrane 2017

## 5.1 Pre-Deterioration Vital Sign Trajectories

| Time Before Event | HR Trend | SpO₂ Trend | Temp Trend | Clinical Significance |
| --- | --- | --- | --- | --- |
| T-60 min | Normal or slight ↑ | Normal | May start rising | Subclinical deterioration begins. Temp rise often the first detectable sign. |
| T-30 min | ↑ 10–20% above baseline | ↓ 1–2% | Rising or peaked | Window where early warning systems should trigger. Ideal intervention point. |
| T-15 min | ↑ 20–30% above baseline | ↓ 2–5% | Peaked or declining | Clear deterioration. NEWS2 score typically 4–6. Urgent review indicated. |
| T-0 (Event) | Tachycardia or bradycardia (terminal) | Below threshold | Variable | Clinical event. NEWS2 ≥7. Emergency response required. |

## 5.2 Vital Sign Relationship Rules for AI Training

These inter-vital relationships encode clinical logic that single-parameter thresholds cannot capture:

| Rule | Pattern | Clinical Interpretation |
| --- | --- | --- |
| Rule 1 | HR rising + SpO₂ falling simultaneously | Respiratory deterioration. More urgent than either sign alone. Likely hypoxia-driven tachycardia. |
| Rule 2 | HR rising + Temp rising simultaneously | Infection/sepsis developing. If both rise >10% from baseline within 30 min: sepsis screen positive. |
| Rule 3 | HR falling + SpO₂ falling simultaneously | CRITICAL — cardiovascular collapse or severe hypoxia. This combination = emergency. |
| Rule 4 | SpO₂ falling + Temp normal | Mechanical hypoxia (airway, PE, pneumothorax). Not infection. Different management pathway. |
| Rule 5 | HR proportional to Temp rise (~10 BPM per 1°C) | Expected fever response. Not an emergency finding if SpO₂ is maintained. |
| Rule 6 | HR elevated + Temp normal + SpO₂ normal | Could be anxiety, pain, dehydration, exertion. Low acuity if all three within tolerable range. |
| Rule 7 | HRV (SDNN) collapsing + SpO₂ dropping | Cardiac involvement. Reduced HRV with SpO₂ drop suggests cardiac-respiratory coupling failure. |
| Rule 8 | Temp >38.3°C + HR >90 (SIRS criteria met) | Sepsis screening positive. Mandatory clinical review. Do not dismiss even if SpO₂ normal. |
| Rule 9 | All three vitals deteriorating simultaneously | Multi-organ involvement. Highest acuity classification. Immediate emergency response. |
| Rule 10 | Rapid SpO₂ drop (>3% in <5 min) | Acute event — not gradual deterioration. Possible airway obstruction, acute bronchospasm, or PE. |

# SECTION 6: SIRS CRITERIA AND SEPSIS SCREENING

Source: Singer M, et al. The Third International Consensus Definitions for Sepsis and Septic Shock (Sepsis-3). JAMA. 2016;315(8):801-810

Full SIRS (Systemic Inflammatory Response Syndrome) requires 2 or more of the following 4 criteria. With only 3 sensors, only 2 of 4 can be assessed:

|  | Full SIRS Criterion | Your Sensor Availability | Implementation |
| --- | --- | --- | --- |
| 1 | Temperature >38.3°C or <36.0°C | LM35 — AVAILABLE | Direct measurement after calibration offset applied |
| 2 | Heart Rate >90 BPM | MAX30102 — AVAILABLE | Direct measurement. Age-adjust: use age-specific upper normal as threshold |
| 3 | Respiratory Rate >20 breaths/min | NOT AVAILABLE | Cannot measure without PPG waveform or dedicated RR sensor. Omit from score. |
| 4 | WBC >12,000 or <4,000 cells/mm³ | NOT AVAILABLE | Requires blood test. Cannot be estimated from wearable sensors. |

## 3-Sensor Sepsis Screening Rule

| Trigger Condition | Definition (3-sensor implementation) |
| --- | --- |
| Sepsis Screen POSITIVE | HR > age-adjusted upper normal (see Section 1) AND Temp > 38.3°C or < 36.0°C — both criteria met simultaneously for ≥2 consecutive readings (≥10 minutes). Output: "Sepsis screening positive — clinical review required." Do NOT output "sepsis diagnosis." |
| Sepsis Screen NEGATIVE | Either HR or Temp criterion not met. Continue monitoring. If only one criterion met, flag as "single SIRS criterion — monitor closely." |
| Important Limitation | This 2-of-2 available criterion screen has lower specificity than full SIRS. Expect higher false positive rate. All positive screens must be verified by clinical assessment. The output is a screening alert, not a diagnosis. |

# SECTION 7: PAEDIATRIC-SPECIFIC CRITICAL THRESHOLDS

Paediatric patients require separate threshold logic. Adult thresholds must never be applied to patients under 18. The following are APLS/PALS validated critical thresholds by age group.

Source: APLS UK Advanced Paediatric Life Support 8th Edition 2021 · AHA PALS Guidelines 2020 · RCPCH

| Age Group | Critical Bradycardia | Critical Tachycardia | SpO₂ Critical | Temp Critical High | Notes |
| --- | --- | --- | --- | --- | --- |
| Neonate (0–28d) | <100 BPM | >180 BPM | <92% | >38.0°C | Any fever in neonate = emergency. Do not wait for HR/SpO₂ changes. |
| Infant (1–12m) | <80 BPM | >180 BPM | <92% | >39.0°C | Fever >39°C in infant under 3 months = emergency without exception. |
| Toddler (1–3y) | <70 BPM | >160 BPM | <92% | >40.0°C | Febrile convulsion risk peaks at this age. Rapid temp rise is as important as absolute value. |
| Preschool (3–5y) | <65 BPM | >150 BPM | <93% | >40.0°C |  |
| School age (6–12y) | <60 BPM | >140 BPM | <93% | >40.0°C | Approach adult thresholds after age 10. |

# SECTION 8: DERIVED PARAMETERS — FORMULAS AND VALID RANGES

The following derived parameters are computable from HR, SpO₂, and Temperature readings alone without raw PPG waveform access. These are the only derived parameters that should be used in model training with MAX30102 processed outputs.

| Parameter | Formula | Valid Range | Clinical Use |
| --- | --- | --- | --- |
| HRV (SDNN) | Std deviation of (60000/HR) over rolling 10-reading window. Requires beat-to-beat BPM from firmware, not smoothed average. | 0–200 ms | Autonomic stress. SDNN <20ms = high stress/acute illness. |
| Metabolic Stress Index | ((HR - baseline_HR)/baseline_HR + (Temp - baseline_Temp)/baseline_Temp) × 50. Baseline = first 5-min resting average. | 0–100 | Infection/inflammation proxy. >30 = moderate stress, >60 = high stress. |
| Temp Rate of Change | (Temp_current - Temp_15min_ago) / 0.25. Unit: °C/hour. | ±10 °C/hr | Rising fever trajectory. >2°C/hr = rapidly escalating fever. |
| SpO₂ Desaturation Flag | TRUE if SpO₂ drops ≥3% from 5-reading moving average AND sustained for ≥2 consecutive readings. | Boolean | Clinically significant desaturation vs sensor noise. |
| Deterioration Trend | Linear regression slope (polyfit) over last 6 readings for HR and SpO₂. Deteriorating = HR slope >2 AND SpO₂ slope <-0.5. | Categorical: improving / stable / deteriorating | Primary AI input for predictive alert generation. |
| Partial NEWS2 Score | Sum of HR score + SpO₂ score + Temp score from Section 3. Maximum possible = 8 (3+3+2). Label output as "Partial NEWS2." | 0–8 | Clinical severity index. Comparable to full NEWS2 if offset for missing RR. |
| SIRS 2-Parameter Check | TRUE if HR > age_upper_normal AND (Temp > 38.3 OR Temp < 36.0). | Boolean | Sepsis screening flag. Positive = clinical review required. |
| O₂-Temp Composite Index | (100 - SpO₂) × (Temp - 37.0). Positive value = hypoxia with fever. Higher = more concerning combined state. | 0 to 200 | Simultaneous hypoxia + fever = pneumonia/sepsis pattern. |

Source: HRV: Task Force of the European Society of Cardiology 1996 · SDNN clinical significance: Kleiger et al., Am J Cardiol 1987

# SECTION 9: INDIA-SPECIFIC POPULATION ADJUSTMENTS

Indian populations have clinically validated differences in metabolic and cardiovascular risk profiles compared to Western populations. Models trained exclusively on MIMIC-IV (US patients) will underperform on Indian patients without the following adjustments.

Source: ICMR Guidelines · WHO Asian BMI Cutoffs · Mohan V, JAPI 2011 · Anand SS, Lancet 2008 · NFHS-5 2019–2021

## 9.1 BMI Cutoff Differences

| Category | WHO Western Cutoff | WHO Asian / Indian Cutoff (Use This) |
| --- | --- | --- |
| Normal | 18.5–24.9 | 18.5–22.9 (lower upper bound) |
| Overweight | 25.0–29.9 | 23.0–27.4 (risk begins earlier) |
| Obese | ≥30.0 | ≥27.5 (lower obesity threshold) |

## 9.2 Disease Prevalence Differences Affecting Baseline Vitals

| Condition | Indian Prevalence | Model Adjustment Required |
| --- | --- | --- |
| Type 2 Diabetes | ~11.4% (ICMR-INDIAB 2023) | Higher proportion of patients will have masked fever response and autonomic neuropathy affecting HR variability. Diabetes comorbidity weight should be higher than in Western models. |
| Hypertension | ~28.5% adults (NFHS-5) | Significant proportion on beta-blockers — HR tachycardia threshold should be adjusted down for this subgroup. Ensure comorbidity flag is used. |
| COPD | ~4.2% (BOLD Study India) | Biomass fuel exposure is primary cause in rural India (not smoking). SpO₂ baseline may be lower in this subgroup. |
| Undernutrition | ~14.6% (NFHS-5) | Low BMI (<18.5) is common in Indian elderly and rural populations. Hypothermia (Temp <36°C) is more likely as a baseline. Adjust hypothermia alert threshold accordingly. |

# SECTION 10: TRAINING DATA SCHEMA AND FIELD DEFINITIONS

The following defines the complete expected schema for each training record. Every field is described with its data type, valid range, and the clinical rule it supports.

| Field Name | Type | Valid Range | Source | Notes |
| --- | --- | --- | --- | --- |
| timestamp | ISO 8601 | 2020 – present | NTP / Firebase | Unix epoch preferred. millis() from ESP32 is NOT a valid timestamp. |
| patient_id | String | Any | System | Anonymised ID. Never store name or Aadhaar in the same table. |
| age | Integer | 0 – 120 | Patient record | In years. Use 0 for neonates (add age_days field for <1 year). |
| age_group | Categorical | See Section 1 | Derived | Must be derived from age at record creation, not stored separately. |
| gender | Categorical | M / F / Other | Patient record | Used for BMI interpretation. Limited vital sign impact. |
| bmi | Float | 10.0 – 70.0 | Calculated | Use Asian BMI cutoffs (Section 9). Flag <18.5 as underweight. |
| comorbidities | Array/String | See Section 2 | Patient record | Comma-separated. Applied as adjustments to all threshold checks. |
| activity_level | Categorical | sedentary / light / moderate / active | Sensor / self-report | HR interpretation depends heavily on activity. Active patient HR 120 BPM is not tachycardia. |
| heartRate | Integer | 20 – 250 BPM | MAX30102 | Reject values outside 20–250 as sensor artifacts. Flag hrValid=false. |
| spO2 | Float | 70.0 – 100.0% | MAX30102 | Reject <70% as likely artifact. Flag spo2Valid=false. Round to 1 decimal. |
| temperature | Float | 30.0 – 44.0°C | LM35 | Apply calibration offset (Section hardware notes). Reject outside 30–44°C. |
| hrv_sdnn | Float | 0 – 200 ms | Derived | Only valid if firmware reports beat-to-beat BPM. Set null otherwise. |
| metabolic_stress | Float | 0 – 100 | Derived | Requires personal baseline from first 5 minutes of session. |
| news2_partial | Integer | 0 – 8 | Derived | Sum of HR + SpO₂ + Temp NEWS2 scores. Max 8 (RR absent). |
| sirs_flag | Boolean | True / False | Derived | True if both 3-sensor SIRS criteria met simultaneously. |
| deterioration_trend | Categorical | improving / stable / deteriorating | Derived | From polyfit over last 6 readings. Primary AI prediction feature. |
| scenario / label | Categorical | See Section 4 | Clinical label | Target variable for supervised learning. Must be assigned by clinician, not algorithm. |

# SECTION 11: DATA QUALITY RULES AND REJECTION CRITERIA

These rules must be applied before any data point enters the training pipeline or triggers a clinical alert. Bad data is worse than no data in a medical AI system.

| Rejection Rule | Threshold | Reason |
| --- | --- | --- |
| HR out of physiological range | HR < 20 or > 250 BPM | Values outside this range are physiologically impossible and indicate sensor artifact or detached sensor. |
| SpO₂ out of range | SpO₂ < 70% or > 100% | SpO₂ <70% is non-survivable for extended periods and almost always a sensor artifact. |
| Temperature out of range | Temp < 30°C or > 44°C | Outside human survival limits. Indicates disconnected or misread LM35. |
| Sensor validity flag false | hrValid=0 or spo2Valid=0 | MAX30102 maxim algorithm reports invalid when signal quality is insufficient. Always check these flags. |
| No finger detected | IR value < 50,000 | IR below 50,000 means no finger on sensor. Reject entire reading. |
| Temperature not stabilised | < 60 seconds since finger detected | LM35 requires thermal equilibration time. Readings before 60 seconds are unreliable. |
| Timestamp invalid | Timestamp = millis() value | millis() resets on reboot. Only Unix epoch NTP timestamps are valid for time-series analysis. |
| Implausible simultaneous values | HR > 200 AND SpO₂ > 99% AND Temp < 36°C | Physiologically implausible combination suggests multi-sensor artifact. Flag for review. |

## 11.1 Minimum Data Requirements for Model Input

- Minimum 3 consecutive valid readings before any alert is generated

- Minimum 10 readings (approximately 90 seconds at 10-second intervals) before HRV SDNN is computed

- Minimum 5-minute resting baseline before Metabolic Stress Index is computed

- At least 6 readings required for deterioration trend calculation

- All three vitals must be valid in the same reading for NEWS2 partial score computation

# SECTION 12: COMPLETE REFERENCE LIST

- Royal College of Physicians. National Early Warning Score (NEWS) 2: Standardising the assessment of acute-illness severity in the NHS. RCP, London, 2017.

- Singer M, et al. The Third International Consensus Definitions for Sepsis and Septic Shock (Sepsis-3). JAMA. 2016;315(8):801–810.

- Global Initiative for Chronic Obstructive Lung Disease (GOLD). Global Strategy for the Diagnosis, Management, and Prevention of COPD. 2023 Report.

- Global Initiative for Asthma (GINA). Global Strategy for Asthma Management and Prevention. 2023 Update.

- Evans L, et al. Surviving Sepsis Campaign: International Guidelines for Management of Sepsis and Septic Shock 2021. Intensive Care Med. 2021;47:1181–1247.

- American Heart Association / American College of Cardiology. 2022 AHA/ACC Guideline for the Diagnosis and Management of Patients with Stable Coronary Artery Disease.

- McDonagh TA, et al. 2021 ESC Guidelines for the Diagnosis and Treatment of Acute and Chronic Heart Failure. Eur Heart J. 2021;42(36):3599–3726.

- American Diabetes Association. Standards of Medical Care in Diabetes 2023. Diabetes Care. 2023;46(Suppl 1).

- Wedzicha JA, et al. Management of COPD Exacerbations: A European Respiratory Society/American Thoracic Society Guideline. Eur Respir J. 2017;49(3).

- British Thoracic Society. Guidelines for the Management of Community Acquired Pneumonia in Adults. BTS, 2023.

- World Health Organization. Pulse Oximetry Training Manual. WHO Press, Geneva, 2011.

- NICE. Fever in under 5s: Assessment and Initial Management. Clinical Guideline CG160. NICE, London, 2021.

- Advanced Paediatric Life Support Group. Advanced Paediatric Life Support: A Practical Approach to Emergencies. 8th Edition. Wiley-Blackwell, 2021.

- American Heart Association. Pediatric Advanced Life Support (PALS) Provider Manual. 2020.

- Task Force of the European Society of Cardiology and the North American Society of Pacing and Electrophysiology. Heart Rate Variability: Standards of Measurement. Circulation. 1996;93:1043–1065.

- Kleiger RE, et al. Decreased Heart Rate Variability and Its Association with Increased Mortality after Acute Myocardial Infarction. Am J Cardiol. 1987;59(4):256–262.

- WHO Expert Consultation. Appropriate Body-Mass Index for Asian Populations and Its Implications for Policy and Intervention Strategies. Lancet. 2004;363:157–163.

- Anjana RM, et al. Prevalence of Diabetes and Prediabetes in 15 States of India: Results from the ICMR-INDIAB Population-Based Cross-Sectional Study. Lancet Diabetes Endocrinol. 2023.

- National Family Health Survey (NFHS-5) 2019–2021. Ministry of Health and Family Welfare, Government of India.

- Churpek MM, et al. Multicenter Comparison of Machine Learning Methods and Conventional Regression for Predicting Clinical Deterioration on the Wards. Crit Care Med. 2016;44(2):368–374.

- India Digital Personal Data Protection Act (DPDPA) 2023. Ministry of Electronics and Information Technology, Government of India.

- Central Drugs Standard Control Organisation (CDSCO). Medical Devices Rules 2017. Ministry of Health and Family Welfare, India.